## Slide 1
TCT characterization and scanning of fine pitch of n+-in-p pixel detectors
http://lnr.irb.hr/PaRaDeSEC/
http://research.hip.fi/hwp/cmsupg/
https://www.aalto.fi/en/otanano
http://www.micronova.fi/
Spatially resolved Charge Collection Efficiency (CCE) of a
pixel detector. The signal is created by 2 MeV proton beam
M. Bezak 1), S. Bharthuar 2), V. Chmill 1), E. J. Brücken 2), A. Gädda 2,3), J.Härkönen 1), M. Kalliokoski 5), A. Karadzhinova-Ferrer 1), S. Kirschemann 2), J. Ott 2),
V. Litichevskyi 2,4), P. Luukka 2)
1) Ruđer Bošković Institute, Bijenička cesta 54, 10000 Zagreb, Croatia
2) Helsinki Institute of Physics, Finland
3) Advacam Oy, Tietotie 3, FI-02150 Espoo, Finland
4) Specom Oy, Tekniikantie 2 A 326, 02150 Espoo, Finland
5) Aalto Univesity, Espoo, Finland
contact: name + @helsinki.fi or @irb.hr

## Slide 2
Outline
Pixel detector design and layout
    AC-coupled pixel sensor, 100×150 µm pitch to match PSI46dig
    DC-coupled pixel sensor, 50×50 µm pitch to match RD53A
Processing of detectors
TCT characterization and selected results
    Image Processing
    High resolution area scans
Summary

## Slide 3
1st level – Implant
2nd level – contact opening
to field insulator
3rd level – bias resistor
made of TiN deposited by Atomic
Layer Deposition (ALD) method
4th level – metals
Detector design – RD53 pixel detector
RD53 = CERN joint collaboration developing future
CMOS read-out ASIC for all LHC experiments. Chip consist
of 400 x 192 = 76 800 pixels

## Slide 4
Detector design – RD53 pixel detector
Presented at 33rd RD50 Workshop, CERN, Geneva, Switzerland, November 28th 2018
J. Ott et al., Processing of pixel detectors on p-type MCz silicon using atomic layer deposition (ALD) grown aluminium oxide
https://indico.cern.ch/event/754063/contributions/3222806/attachments/1760772/2865963/JOtt_RD50_Nov18_3.pdf
38µm

## Slide 5
Detector design – PSI46dig pixel detector
Presented at 33rd RD50 Workshop, CERN, Geneva, Switzerland, November 28th 2018
J. Ott et al., Processing of pixel detectors on p-type MCz silicon using atomic layer deposition (ALD) grown aluminium oxide
https://indico.cern.ch/event/754063/contributions/3222806/attachments/1760772/2865963/JOtt_RD50_Nov18_3.pdf
70µm
120µm

## Slide 6
Processing of detectors

## Slide 7
Characterization and selected results

## Slide 8
Raster scanning surface TCT setup from Particulars d.o.o
Currently with two wavelengths 660 nm (red) and 1064 nm (IR)
Readout with 4 GHz Lecroy Waverunner oscilloscope
TCT-setup

## Slide 9
TCT- Apsolute Calibration of Amplifier
Asymetrical behavior
  1000 mV – positive signal
  600 mV – negative signal

## Slide 10
TCT- Apsolute Calibration of Amplifier
Gain of amplifier (310 ± 3) which correspondence to 50 dB
Important for translation to MIP equvivalent
* 53 dB ≈ 447 gain

## Slide 11
TCT Voltage Scan
Full depletion voltage Vfd

## Slide 12
After the optical focal distance is determined, XY-scan of the selected detector areas is made
TCT Area Scans
RD53 pixel detector
660nm
1064nm
100 V

## Slide 13
Pixel Statistics
Homogeneity of scanned pixels in terms of collected charge
Reference point – max value of collected charge

## Slide 14
TCT Area Scans for different bias Voltages
1064nm

## Slide 15
TCT Area Scans for different bias Voltages
660nm

## Slide 16
TCT Big Area Scans
1064nm

## Slide 17
Big Area Scans
By performing measurments impact of mechanical (stage) movment was observed
Solutions :
  Increasing the waiting time while thaking the wave forms
  Applaying procesing tools - Discrete Fourier Transformat
Figure 21 360 X 10 960 μm
Waiting time t = 0.1 s
Step dr=(15,15, 0) μm
Number of points N=(1425, 548, 0)
Time of scaning 780900 points
      = 21h41min24s

## Slide 18
Big Area Scans - DFT
2 dim. DFT -> Magnitude Spectrum
The Fourier image is shifted in such a way that the DC-value (image mean) F(0,0) is in the center
Image contains components of all frequencies – magnitude is smaller for higher frequencies
Dominating direction – background pattern
2D DFT + SHIFT

## Slide 19
Big Area Scans - DFT
Bright spikes
Grey Scale
Magnitude Spectrum
Noise image
Inverse FFT
Coloring*

## Slide 20
660nm
1064nm

## Slide 21
Summary
  Helsinki Institute of Physics (HIP) and RBI PaRaDeSeC teams have on-going common activity to develop new and innovative pixel detectors intented for CMS Phase II Upgrade
  We have simulated, designed, manufactured and characterized novel fine pitch n+/p-/p+ pixel detectors made on 150mm size p-type Magnetic Czochralski silicon (p-MCz Si) wafers
  Detectors are processed at Micronova nanofabrication
  center in Finland
  Atomic Layer Deposition (ALD) technology has many properties, which make it very attractive process method for radiation detectors.
  TCT can be used to study the uniformity and possible radiation damage
  By studiying image processing tools more detailed image of detector performance can be done

## Slide 22
Backup slides
When pixel technology meets Arts....
Each square is TCT xy-scan of about
60 000 data points.
The amplitude of the signal is converted
into color scale.

## Slide 23
Backup slides
-Micronova facility

## Slide 24
Pixel Statistics

## Slide 25
Pixel Statistics

## Slide 26
Big Area Scans

## Slide 27
Passivation – Atomic Layer Deposition (ALD)
J. Härkönen et al., ”Processing of n+/p−/p+ strip detectors with atomic layer deposition (ALD) grown Al2O3 field insulator on magnetic Czochralski silicon (MCz-si) substrates”, NIMA 826, 2016
J. Härkönen et al., Atomic Layer Deposition (ALD) grown thin films for ultra-fine pitch pixel detectors, NIMA 831, 2016
J. Ott et al., Processing of AC-coupled n-in-p pixel detectors on MCz silicon using atomic layer deposited aluminium oxide, in VCI 2019 - The 15th Vienna Conference on Instrumentation, February 20, 2019
The electrical passivation properties
can be studied by lifetime (τeff) measurement
τeff is combination of bulk lifetime (τbulk)
and surface recombination (Ssrv).
High τeff can only be measured if Ssrv
→ 0, i.e. Si surface is passivated.
Thermally oxidized Si-SiO2 interface is known
to produce best possible Ssrv→ 0
Oxidized p-type reveals bulk lifetime (τbulk)
(BLUE distribution) and thus it is reference value for passivation studies.
SiO2 is removed  ALD Al2O3 deposition
+ repeated lifetime measurement (RED distribution).
Good passivation (Ssrv→ 0) is achieved by field effect, negative oxide
oxide charge in Al2O3 is repulsing e- to recombinate into surface states.

## Slide 28
Full depletion voltage Vfd

## Slide 29
Acknowledgements
This research has recived funding from the European Union’s Horizon 2020 ERA Chair reasearch and innovation programme under grant agreement No. 669014, PaRaDeSEC
Micronova Nanofabrication Centre